## Slide 1: EZRC - LN2 containers in Bldg. 341
2 mobile storage tanks are parked next to side entrance
Refilled weekly by delivery service
4 sample containers should be refilled twice weekly as shown below
Open this door to move
Check brake is not set
Check that tube is attached to “LIQUID” outlet
Only operate blue valve

## Slide 2: EZRC - LN2 containers in Bldg. 341 hallway
Keys and gloves and (if needed) wrenches are in a “secret” compartment in the cryogenic freezer
Pull lever above display to open

## Slide 3: EZRC - LN2 containers in Bldg. 341 hallway
Fill square containers until towers incl. handles are just covered

## Slide 4: EZRC - LN2 containers in Bldg. 341 hallway
One key opens both round containers
Fill until liquid just reaches the neck
Max. liquid level

## Slide 5: EZRC - LN2 containers in Bldg. 341 hallway
Lock like this
Sign list

## Slide 6: EZRC – cryogenic freezer in Bldg. 341 hallway
Emergency LN2 tank is refilled and reattached weekly by delivery service
Freezer is automatically flooded if temperature rises
Valve to freezer must be open at all times
In emergency keep filled with LN2 until serviced by manufacturer
-147 to -150 °C is normal, must never exceed -130 °C

## Slide 7: EZRC - 80 °C freezers in Bldg. 341 hallway
Contain samples from Tübingen screens, not high priority
In case of break-down empty out and discard as there is very little reserve space (garbage bags in cleaning kitchen next to side entrance)
Save boxes labelled “SC” or “FDA library”, they belong to the Screening Center
Also ”ZF PAC library”

## Slide 8: EZRC – 2 cryogenic freezers in Bldg. 223 garage
Emergency LN2 tanks are refilled and reattached weekly by delivery service
Freezers are automatically flooded if temperature rises
In emergency keep filled with LN2 until serviced by manufacturer
Keys hidden here
Door is open

## Slide 9: EZRC – 2 cryogenic freezers in Bldg. 223 garage
Emergency LN2 tanks should be attached like this (normally by delivery service)
Valves to freezer must be open at all times

## Slide 10: EZRC – Bldg. 223 lab
-20 °C
-20 °C reserve
Refrigerator
-20 °C
-20 °C, some reserve space
-20 °C for dead fish

## Slide 11: EZRC – Bldg. 223 office
code 673425
fish food refrigerator
- 20 °C, some reserve space

## Slide 12: EZRC – 2 cryogenic freezers in Bldg. 223 garage
Emergency LN2 tanks are refilled and reattached weekly by delivery service
Freezers are automatically flooded if temperature rises
In emergency keep filled with LN2 until serviced by manufacturer
Keys hidden here
Door is open

## Slide 13: EZRC – 2 cryogenic freezers in Bldg. 223 garage
Emergency LN2 tanks should be attached like this (normally by delivery service)
Valves to freezer must be open at all times

## Slide 14: EZRC – Bldg. 223 lab
-20 °C
-20 °C reserve
Refrigerator
-20 °C
-20 °C, some reserve space
-20 °C for dead fish

## Slide 15: EZRC – Bldg. 223 office
code 673425
fish food refrigerator
- 20 °C, some reserve space

## Slide 16
Perkin Elmer Devices 341
Not to be switched off!

## Slide 17
Olympus Devices 341
The computers are not to be switched off! The light sources should be switched off after the experiments.

## Slide 18
NGS 341